## Manual Flow Measurements
Site Name: San Dieguito Creek
Pipe Diameter (in) | 24
Flow Measurement Number | Datetime | Raw Level from PT (in) | Manual Flow Measurement (cfs) | Sutron Flow (cfs) | ABSOLUTE ERROR (cfs) | RELATIVE ERROR (%) | Water Depth (ft) | Invert (ft) | Field Staff | Comments
1 | 2020-03-26 15:00 | 2.79 | 0.988 | 1.8441 | -0.85613 | -46.42% | N/A | N/A | AM
2 | 2020-03-31 10:20 | 2.67 | 1.8647 | 1.6919 | 0.17284 | 10.22% | N/A | N/A | JE
3 | 2020-03-12 14:00 | 0.04 | 0 | 0.013784 | -0.013784 | -100% | 0.019167 | 0.0625 | AM | Omitted
4 | 2020-04-07 09:10 | 5.73 | 4.4089 | 7.5035 | -3.0946 | -41.24% | N/A | N/A | AM
5 | 2020-04-08 15:15 | 10.03 | 12.8269 | 8.2937 | 4.5332 | 54.66% | N/A | N/A | JE
6 | 2020-04-09 11:20 | 38.67 | 87.1792 | 765.4118 | -678.2326 | -88.61% | N/A | N/A | AM
7 | 2020-04-09 12:05 | 39.79 | 97.3894 | 807.4052 | -710.0159 | -87.94% | N/A | N/A | AM
8 | 2020-06-11 15:23 | 1.66 | 0.003984 | 0.67363 | -0.66965 | -99.41% | N/A | N/A | JE JN
9 | 2020-06-11 15:37 | 1.57 | 0.0030412 | 0.59889 | -0.59585 | -99.49% | N/A | N/A | JE JN
10 | 2020-06-11 15:46 | 1.55 | 0.0041969 | 0.58228 | -0.57808 | -99.28% | N/A | N/A | JE JN
11 | 2020-06-16 14:45 | 1.4 | 0.0039077 | 0.48245 | -0.47854 | -99.19% | N/A | N/A | JE JN
12 | 2020-06-16 14:50 | 1.39 | 0.0035294 | 0.479 | -0.47547 | -99.26% | N/A | N/A | JE JN
13 | 2020-06-26 11:40 | 0.88 | 0.0021551 | 0.30325 | -0.3011 | -99.29% | N/A | N/A | JE KG
14 | 2020-06-26 11:50 | 0.87 | 0.0023859 | 0.29981 | -0.29742 | -99.2% | N/A | N/A | JE KG
|  | Length | t1 | t2 | t3 | Average Velocity (ft/s) | Depth above invert (in) | Angle | Area,ft2 | Flow (cfs)
Flow Measurement 2 | Pipe 1 | 20 | 11 | 12.93 | 12.8 | 1.6425 | 3.5 | 89.8011 | 0.28367 | 0.46592
| Pipe 2 | 20 | 11.63 | 11.63 | 12 | 1.702 | 3.75 | 93.1349 | 0.3135 | 0.53359
| Pipe 3 | 20 | 15.14 | 12.74 | 13 | 1.4764 | 4.25 | 99.5436 | 0.3756 | 0.55455
| Pipe 4 | 20 | 13.18 | 11.96 | 13.52 | 1.5563 | 2.75 | 79.1424 | 0.1996 | 0.31064
[9 empty] | Total Flow (cfs) | 1.8647
|  | Length | t1 | t2 | t3 | Average Velocity (ft/s) | Depth above invert (in) | Angle | Area,ft2 | Flow (cfs)
Flow Measurement 4 | Pipe 1 | 20 | 8 | NR | NR | 2.5 | 4.75 | 105.6622 | 0.44064 | 1.1016
| Pipe 2 | 20 | NR | NR | NR | NR | NR | NR | NR | 1.2616
| Pipe 3 | 20 | NR | NR | NR | NR | NR | NR | NR | 1.3112
| Pipe 4 | 20 | NR | NR | NR | NR | NR | NR | NR | 0.73447
[9 empty] | Total Flow (cfs) | 4.4089
|  | Inlet Velocity (ft/s) | Outlet Velocity (ft/s) | Average Velocity (ft/s) |  |  | Depth above invert (ft) | Angle | Area (ft2) | Flow (cfs)
Flow Measurement 5 | Pipe 1 | 1.55 | 3.3 | 2.425 |  |  | 10.2 | 162.7461 | 1.2719 | 3.0844
| Pipe 2 | 1.26 | 3.38 | 2.32 |  |  | 10.8 | 168.5217 | 1.3711 | 3.181
| Pipe 3 | 1.24 | 3.09 | 2.165 |  |  | 12 | 180 | 1.5708 | 3.4008
| Pipe 4 | 2.07 | 2.9 | 2.485 |  |  | 10.2 | 162.7461 | 1.2719 | 3.1607
[9 empty] | Total Flow (cfs) | 12.8269
|  | Inlet Velocity (ft/s) | Outlet Velocity (ft/s) | Average Velocity (ft/s) |  |  | Depth above invert (ft) | Angle | Area (ft2) | Flow (cfs)
Flow Measurement 6 | Pipe 1 | NR | NR | 7.5 |  |  | 24 | 0 | 3.1416 | 23.5619
| Pipe 2 | NR | NR | 8.25 |  |  | 24 | 0 | 3.1416 | 25.9181
| Pipe 3 | NR | NR | 6.5 |  |  | 24 | 0 | 3.1416 | 20.4204
| Pipe 4 | NR | NR | 5.5 |  |  | 24 | 0 | 3.1416 | 17.2788
[9 empty] | Total Flow (cfs) | 87.1792
|  | Inlet Velocity (ft/s) | Outlet Velocity (ft/s) | Average Velocity (ft/s) |  |  | Depth above invert (ft) | Angle | Area (ft2) | Flow (cfs)
Flow Measurement 7 | Pipe 1 | NR | NR | 7.5 |  |  | 24 | 0 | 3.1416 | 23.5619
| Pipe 2 | NR | NR | 8.5 |  |  | 24 | 0 | 3.1416 | 26.7035
| Pipe 3 | NR | NR | 7.8 |  |  | 24 | 0 | 3.1416 | 24.5044
| Pipe 4 | NR | NR | 7.2 |  |  | 24 | 0 | 3.1416 | 22.6195
[9 empty] | Total Flow (cfs) | 97.3894
|  | Inlet Velocity (ft/s) | Outlet Velocity (ft/s) | Average Velocity (ft/s) |  |  | Depth above invert (ft) | Angle | Area (ft2) | Flow (cfs)
Flow Measurement 8 | Pipe 1 | NR | NR | 0.28 |  |  | 0.2 | 20.9506 | 0.004047 | 0.0011332
| Pipe 2 | NR | NR | 0.15 |  |  | 0.425 | 30.5887 | 0.012501 | 0.0018751
| Pipe 3 | NR | NR | 0.1 |  |  | 0.275 | 24.5797 | 0.006519 | 0.0006519
| Pipe 4 | NR | NR | 0.08 |  |  | 0.2 | 20.9506 | 0.004047 | 0.00032376
[9 empty] | Total Flow (cfs) | 0.003984
|  | Inlet Velocity (ft/s) | Outlet Velocity (ft/s) | Average Velocity (ft/s) |  |  | Depth above invert (ft) | Angle | Area (ft2) | Flow (cfs)
Flow Measurement 9 | Pipe 1 | NR | NR | 0.23 |  |  | 0.2 | 20.9506 | 0.004047 | 0.00093082
| Pipe 2 | NR | NR | 0.11 |  |  | 0.42 | 30.4072 | 0.012282 | 0.001351
| Pipe 3 | NR | NR | 0.06 |  |  | 0.275 | 24.5797 | 0.006519 | 0.00039114
| Pipe 4 | NR | NR | 0.14 |  |  | 0.15 | 18.1374 | 0.0026303 | 0.00036824
[9 empty] | Total Flow (cfs) | 0.0030412
|  | Inlet Velocity (ft/s) | Outlet Velocity (ft/s) | Average Velocity (ft/s) |  |  | Depth above invert (ft) | Angle | Area (ft2) | Flow (cfs)
Flow Measurement 10 | Pipe 1 | NR | NR | 0.21 |  |  | 0.2 | 20.9506 | 0.004047 | 0.00084988
| Pipe 2 | NR | NR | 0.14 |  |  | 0.42 | 30.4072 | 0.012282 | 0.0017195
| Pipe 3 | NR | NR | 0.12 |  |  | 0.3 | 25.6771 | 0.0074255 | 0.00089107
| Pipe 4 | NR | NR | 0.28 |  |  | 0.15 | 18.1374 | 0.0026303 | 0.00073648
[9 empty] | Total Flow (cfs) | 0.0041969
|  | Inlet Velocity (ft/s) | Outlet Velocity (ft/s) | Average Velocity (ft/s) |  |  | Depth above invert (ft) | Angle | Area (ft2) | Flow (cfs)
Flow Measurement 11 | Pipe 1 | NR | NR | 0.24 |  |  | 0.175 | 19.5941 | 0.0033135 | 0.00079524
| Pipe 2 | NR | NR | 0.16 |  |  | 0.4 | 29.6702 | 0.011418 | 0.0018269
| Pipe 3 | NR | NR | 0.12 |  |  | 0.3 | 25.6771 | 0.0074255 | 0.00089107
| Pipe 4 | NR | NR | 0.15 |  |  | 0.15 | 18.1374 | 0.0026303 | 0.00039454
[9 empty] | Total Flow (cfs) | 0.0039077
|  | Inlet Velocity (ft/s) | Outlet Velocity (ft/s) | Average Velocity (ft/s) |  |  | Depth above invert (ft) | Angle | Area (ft2) | Flow (cfs)
Flow Measurement 12 | Pipe 1 | NR | NR | 0.16 |  |  | 0.175 | 19.5941 | 0.0033135 | 0.00053016
| Pipe 2 | NR | NR | 0.16 |  |  | 0.4 | 29.6702 | 0.011418 | 0.0018269
| Pipe 3 | NR | NR | 0.1 |  |  | 0.3 | 25.6771 | 0.0074255 | 0.00074255
| Pipe 4 | NR | NR | 0.3 |  |  | 0.1 | 14.804 | 0.0014326 | 0.00042979
[9 empty] | Total Flow (cfs) | 0.0035294
|  | Inlet Velocity (ft/s) | Outlet Velocity (ft/s) | Average Velocity (ft/s) |  |  | Depth above invert (ft) | Angle | Area (ft2) | Flow (cfs)
Flow Measurement 13 | Pipe 1 | NR | NR | 0.15 |  |  | 0.14 | 17.5212 | 0.002372 | 0.0003558
| Pipe 2 | NR | NR | 0.05 |  |  | 0.4 | 29.6702 | 0.011418 | 0.0005709
| Pipe 3 | NR | NR | 0.09 |  |  | 0.35 | 27.7442 | 0.0093514 | 0.00084162
| Pipe 4 | NR | NR | 0.27 |  |  | 0.1 | 14.804 | 0.0014326 | 0.00038681
[9 empty] | Total Flow (cfs) | 0.0021551
|  | Inlet Velocity (ft/s) | Outlet Velocity (ft/s) | Average Velocity (ft/s) |  |  | Depth above invert (ft) | Angle | Area (ft2) | Flow (cfs)
Flow Measurement 14 | Pipe 1 | NR | NR | 0.17 |  |  | 0.14 | 17.5212 | 0.002372 | 0.00040324
| Pipe 2 | NR | NR | 0.1 |  |  | 0.4 | 29.6702 | 0.011418 | 0.0011418
| Pipe 3 | NR | NR | 0.07 |  |  | 0.35 | 27.7442 | 0.0093514 | 0.0006546
| Pipe 4 | NR | NR | 0.13 |  |  | 0.1 | 14.804 | 0.0014326 | 0.00018624
[9 empty] | Total Flow (cfs) | 0.0023859
Note on K35: Administrator: I made some simplifying assumptions since we only had one leaf float measurement for one pipe. Assumed the velocities/flows in each pipe would be distributed as they were on 3/31/2020. This checks out either way since 4*1.1=4.4.

## HvF table
HEC-RAS rating curve |  | Calibrated Rating Curve
Stage (in) | Q Total (cfs) | Calibrated Stage (in) | Calibrated Flow (cfs)
0 | 0 | 0 | 0
1.4509 | 0.5 | 16 | 15
2.053 | 1 | 17.4509 | 15.5
2.5188 | 1.5 | 18.053 | 16
2.9128 | 2 | 18.5188 | 16.5
3.2607 | 2.5 | 18.9128 | 17
3.576 | 3 | 19.2607 | 17.5
3.8672 | 3.5 | 19.576 | 18
4.1382 | 4 | 19.8672 | 18.5
4.3938 | 4.5 | 20.1382 | 19
4.6355 | 5 | 20.3938 | 19.5
4.8662 | 5.5 | 20.6355 | 20
5.087 | 6 | 20.8662 | 20.5
5.2991 | 6.5 | 21.087 | 21
5.5038 | 7 | 21.2991 | 21.5
5.7012 | 7.5 | 21.5038 | 22
9.8452 | 8 | 21.7012 | 22.5
10.1598 | 8.5 | 25.8452 | 23
10.4498 | 9 | 26.1598 | 23.5
10.7153 | 9.5 | 26.4498 | 24
10.9629 | 10 | 26.7153 | 24.5
11.2101 | 10.5 | 26.9629 | 25
11.4382 | 11 | 27.2101 | 25.5
11.6503 | 11.5 | 27.4382 | 26
11.859 | 12 | 27.6503 | 26.5
12.0582 | 12.5 | 27.859 | 27
12.2509 | 13 | 28.0582 | 27.5
12.4365 | 13.5 | 28.2509 | 28
12.6167 | 14 | 28.4365 | 28.5
12.8071 | 14.5 | 28.6167 | 29
12.9562 | 15 | 28.8071 | 29.5
13.1191 | 15.5 | 28.9562 | 30
13.277 | 16 | 29.1191 | 30.5
13.4634 | 16.5 | 29.277 | 31
13.6084 | 17 | 29.4634 | 31.5
13.7688 | 17.5 | 29.6084 | 32
13.8578 | 18 | 29.7688 | 32.5
14.0372 | 18.5 | 29.8578 | 33
14.1493 | 19 | 30.0372 | 33.5
14.2994 | 19.5 | 30.1493 | 34
14.4492 | 20 | 30.2994 | 34.5
14.7235 | 21 | 30.4492 | 35
14.9235 | 22 | 30.7235 | 36
15.1589 | 23 | 30.9235 | 37
15.3907 | 24 | 31.1589 | 38
15.5651 | 25 | 31.3907 | 39
15.8793 | 26 | 31.5651 | 40
16.0865 | 27 | 31.8793 | 41
16.2894 | 28 | 32.0865 | 42
16.4879 | 29 | 32.2894 | 43
16.6816 | 30 | 32.4879 | 44
16.8721 | 31 | 32.6816 | 45
17.0592 | 32 | 32.8721 | 46
17.2438 | 33 | 33.0592 | 47
17.4243 | 34 | 33.2438 | 48
17.5939 | 35 | 33.4243 | 49
17.785 | 36 | 33.5939 | 50
17.952 | 37 | 33.785 | 51
18.1212 | 38 | 33.952 | 52
18.2867 | 39 | 34.1212 | 53
18.4526 | 40 | 34.2867 | 54
18.6207 | 41 | 34.4526 | 55
18.7848 | 42 | 34.6207 | 56
18.9492 | 43 | 34.7848 | 57
19.092 | 44 | 34.9492 | 58
19.2495 | 45 | 35.092 | 59
19.4125 | 46 | 35.2495 | 60
19.5623 | 47 | 35.4125 | 61
19.7106 | 48 | 35.5623 | 62
19.8611 | 49 | 35.7106 | 63
20.0134 | 50 | 35.8611 | 64
20.1515 | 51 | 36.0134 | 65
20.3035 | 52 | 36.1515 | 66
20.443 | 53 | 36.3035 | 67
20.5657 | 54 | 36.443 | 68
20.7378 | 55 | 36.5657 | 69
20.8513 | 56 | 36.7378 | 70
20.996 | 57 | 36.8513 | 71
21.1047 | 58 | 36.996 | 72
21.2856 | 59 | 37.1047 | 73
21.382 | 60 | 37.2856 | 74
21.4801 | 61 | 37.382 | 75
21.6661 | 62 | 37.4801 | 76
21.7701 | 63 | 37.6661 | 77
21.8342 | 64 | 37.7701 | 78
22.0547 | 65 | 37.8342 | 79
22.1667 | 66 | 38.0547 | 80
22.2627 | 67 | 38.1667 | 81
22.3312 | 68 | 38.2627 | 82
22.4348 | 69 | 38.3312 | 83
22.5414 | 70 | 38.4348 | 84
22.6238 | 71 | 38.5414 | 85
22.8593 | 72 | 38.6238 | 86
22.8743 | 73 | 38.8593 | 87
22.8596 | 74 | 38.8743 | 88
22.9655 | 75 | 38.8596 | 89
23.2324 | 76 | 38.9655 | 90
23.2339 | 77 | 39.2324 | 91
23.38 | 78 | 39.2339 | 92
23.3284 | 79 | 39.38 | 93
23.4419 | 80 | 39.3284 | 94
23.644 | 81 | 39.4419 | 95
23.6689 | 82 | 39.644 | 96
23.7407 | 83 | 39.6689 | 97
23.7301 | 84 | 39.7407 | 98
23.9952 | 85 | 39.7301 | 99
23.8282 | 86 | 39.9952 | 100
24.1403 | 87 | 39.8282 | 101
24.2256 | 88 | 40.1403 | 102
24.2959 | 89 | 40.2256 | 103
24.2662 | 90 | 40.2959 | 104
24.3801 | 91 | 40.2662 | 105
24.5259 | 92 | 40.3801 | 106
24.4958 | 93 | 40.5259 | 107
24.3647 | 94 | 40.4958 | 108
24.5405 | 95 | 40.3647 | 109
24.8954 | 96 | 40.5405 | 110
24.8694 | 97 | 40.8954 | 111
24.6888 | 98 | 40.8694 | 112
24.8445 | 99 | 40.6888 | 113
25.0869 | 100 | 40.8445 | 114
25.2931 | 105 | 41.0869 | 115
25.6952 | 110 | 41.2931 | 120
25.782 | 115 | 41.6952 | 125
26.0577 | 120 | 41.782 | 130
26.484 | 125 | 42.0577 | 135
26.688 | 130 | 42.484 | 140
26.8129 | 135 | 42.688 | 145
26.9996 | 140 | 42.8129 | 150
27.4061 | 145 | 42.9996 | 155
27.472 | 150 | 43.4061 | 160
28.4678 | 155 | 43.472 | 165
27.4519 | 160 | 44.4678 | 170
27.5629 | 165 | 43.4519 | 175
28.2308 | 170 | 43.5629 | 180
28.0298 | 175 | 44.2308 | 185
28.214 | 180 | 44.0298 | 190
28.0151 | 185 | 44.214 | 195
28.4897 | 190 | 44.0151 | 200
28.8347 | 195 | 44.4897 | 205
28.5568 | 200 | 44.8347 | 210
29.0804 | 205 | 44.5568 | 215
29.1885 | 210 | 45.0804 | 220
29.6569 | 215 | 45.1885 | 225
29.3822 | 220 | 45.6569 | 230
29.6953 | 225 | 45.3822 | 235
29.6949 | 230 | 45.6953 | 240
29.9451 | 235 | 45.6949 | 245
29.9557 | 240 | 45.9451 | 250
29.7766 | 245 | 45.9557 | 255
29.9022 | 250 | 45.7766 | 260
29.6913 | 255 | 45.9022 | 265
30.405 | 260 | 45.6913 | 270
30.8251 | 265 | 46.405 | 275
30.6581 | 270 | 46.8251 | 280
30.6321 | 275 | 46.6581 | 285
30.6317 | 280 | 46.6321 | 290
30.8145 | 285 | 46.6317 | 295
31.0734 | 290 | 46.8145 | 300
31.1027 | 295 | 47.0734 | 305
31.1733 | 300 | 47.1027 | 310
31.4865 | 310 | 47.1733 | 315
31.7714 | 320 | 47.4865 | 325
31.8926 | 330 | 47.7714 | 335
32.6133 | 340 | 47.8926 | 345
32.3108 | 350 | 48.6133 | 355
33.0784 | 360 | 48.3108 | 365
32.751 | 370 | 49.0784 | 375
33.0234 | 380 | 48.751 | 385
33.1516 | 390 | 49.0234 | 395
33.3864 | 400 | 49.1516 | 405
33.624 | 410 | 49.3864 | 415
33.6639 | 420 | 49.624 | 425
33.913 | 430 | 49.6639 | 435
33.9818 | 440 | 49.913 | 445
34.1777 | 450 | 49.9818 | 455
34.4253 | 460 | 50.1777 | 465
34.5374 | 470 | 50.4253 | 475
34.7054 | 480 | 50.5374 | 485
34.838 | 490 | 50.7054 | 495
35.2815 | 500 | 50.838 | 505
35.1332 | 520 | 51.2815 | 515
35.532 | 540 | 51.1332 | 535
36.0707 | 560 | 51.532 | 555
36.1443 | 580 | 52.0707 | 575
36.4226 | 600 | 52.1443 | 595
36.6892 | 620 | 52.4226 | 615
36.947 | 640 | 52.6892 | 635
37.2799 | 660 | 52.947 | 655
37.5787 | 680 | 53.2799 | 675
37.8204 | 700 | 53.5787 | 695
38.4818 | 720 | 53.8204 | 715
38.5803 | 740 | 54.4818 | 735
38.3276 | 760 | 54.5803 | 755
39.5929 | 780 | 54.3276 | 775
39.657 | 800 | 55.5929 | 795
40.0162 | 820 | 55.657 | 815
39.7906 | 840 | 56.0162 | 835
40.6238 | 860 | 55.7906 | 855
41.0079 | 880 | 56.6238 | 875
41.4119 | 900 | 57.0079 | 895
41.8799 | 920 | 57.4119 | 915
42.1835 | 940 | 57.8799 | 935
41.941 | 960 | 58.1835 | 955
42.7603 | 980 | 57.941 | 975
43.025 | 1000 | 58.7603 | 995
43.7849 | 1050 | 59.025 | 1015
44.239 | 1100 | 59.7849 | 1065
45.3962 | 1150 | 60.239 | 1115
45.8405 | 1200 | 61.3962 | 1165
46.3967 | 1250 | 61.8405 | 1215
47.2383 | 1300 | 62.3967 | 1265
47.8605 | 1350 | 63.2383 | 1315
48.4329 | 1400 | 63.8605 | 1365
49.0408 | 1450 | 64.4329 | 1415
49.5469 | 1500 | 65.0408 | 1465
53.9568 | 2000 | 65.5469 | 1515
59.9996 | 3000 | 69.9568 | 2015
64.9757 | 4000 | 75.9996 | 3015
68.8085 | 5000 | 80.9757 | 4015
72.1791 | 6000 | 84.8085 | 5015
72.6888 | 6210 | 88.1791 | 6015
77.9542 | 8000 | 88.6888 | 6225
82.8611 | 10000 | 93.9542 | 8015
87.5944 | 12000 | 98.8611 | 10015
89.3463 | 12800 | 103.5944 | 12015
91.8724 | 14000 | 105.3463 | 12815
95.7458 | 16000 | 107.8724 | 14015
99.3713 | 18000 | 111.7458 | 16015
102.8276 | 20000 | 115.3713 | 18015
106.1111 | 22000 | 118.8276 | 20015
109.2389 | 24000 | 122.1111 | 22015
111.9335 | 25700 | 125.2389 | 24015
118.0613 | 30000 | 127.9335 | 25715
124.6011 | 35000 | 134.0613 | 30015
130.1517 | 39400 | 140.6011 | 35015
136.845 | 45000 | 146.1517 | 39415
142.3674 | 50000 | 152.845 | 45015
148.9149 | 56300 | 158.3674 | 50015
152.5902 | 60000 | 164.9149 | 56315
157.3773 | 65000 | 168.5902 | 60015
162.0293 | 70000 | 173.3773 | 65015
168.8503 | 77700 | 178.0293 | 70015
179.4261 | 90000 | 184.8503 | 77715
187.2968 | 100000 | 195.4261 | 90015
194.8923 | 110000 | 203.2968 | 100015
|  | 210.8923 | 110015

## Rating Curve Calibration 1
DateTime | PT Stage (in) | Manually Measured Flow (cfs) | Sutron Flow (cfs) | Absolute Error (cfs) | Relative Error (%)
2020-03-26 15:00 | 2.79 | 0.988 | 1.8441 | -0.85613 | 46.4245
2020-03-31 10:20 | 2.67 | 1.8647 | 1.6919 | 0.17284 | 10.2161
2020-04-07 09:10 | 5.73 | 4.4089 | 7.5035 | -3.0946 | 41.2425
2020-04-08 15:15 | 10.03 | 12.8269 | 8.2937 | 4.5332 | 54.6588
2020-04-09 11:20 | 38.67 | 87.1792 | 765.4118 | -678.2326 | 88.6102
2020-04-09 12:05 | 39.79 | 97.3894 | 807.4052 | -710.0159 | 87.938
2020-03-12 14:00 | 0.04 | 0 | 0.013784 | -0.013784 | 100
2020-06-11 15:23 | 1.66 | 0.003984 | 0.67363 | -0.66965 | 99.4086
2020-06-11 15:37 | 1.57 | 0.0030412 | 0.59889 | -0.59585 | 99.4922
2020-06-11 15:46 | 1.55 | 0.0041969 | 0.58228 | -0.57808 | 99.2792
2020-06-16 14:45 | 1.4 | 0.0039077 | 0.48245 | -0.47854 | 99.19
2020-06-16 14:50 | 1.39 | 0.0035294 | 0.479 | -0.47547 | 99.2632
HEC-RAS rating curve
Stage (in) | Q Total (cfs) | Forecast Relative Error (%) | Revised Flow (cfs)
0 | 0 | #NUM! | #NUM!
1.4509 | 0.5 | 12.1717 | 0.43914
2.053 | 1 | 20.1148 | 0.79885
2.5188 | 1.5 | 24.7946 | 1.1281
2.9128 | 2 | 28.1208 | 1.4376
3.2607 | 2.5 | 30.7028 | 1.7324
3.576 | 3 | 32.8152 | 2.0155
3.8672 | 3.5 | 34.6064 | 2.2888
4.1382 | 4 | 36.1564 | 2.5537
4.3938 | 4.5 | 37.528 | 2.8112
4.6355 | 5 | 38.7535 | 3.0623
4.8662 | 5.5 | 39.8651 | 3.3074
5.087 | 6 | 40.8807 | 3.5472
5.2991 | 6.5 | 41.8153 | 3.782
5.5038 | 7 | 42.6827 | 4.0122
5.7012 | 7.5 | 43.4891 | 4.2383
9.8452 | 8 | 55.9915 | 3.5207
10.1598 | 8.5 | 56.7112 | 3.6795
10.4498 | 9 | 57.3554 | 3.838
10.7153 | 9.5 | 57.9296 | 3.9967
10.9629 | 10 | 58.4523 | 4.1548
11.2101 | 10.5 | 58.9626 | 4.3089
11.4382 | 11 | 59.4237 | 4.4634
11.6503 | 11.5 | 59.844 | 4.6179
11.859 | 12 | 60.2504 | 4.77
12.0582 | 12.5 | 60.6317 | 4.921
12.2509 | 13 | 60.9944 | 5.0707
12.4365 | 13.5 | 61.3386 | 5.2193
12.6167 | 14 | 61.6678 | 5.3665
12.8071 | 14.5 | 62.0106 | 5.5085
12.9562 | 15 | 62.2754 | 5.6587
13.1191 | 15.5 | 62.5614 | 5.803
13.277 | 16 | 62.8351 | 5.9464
13.4634 | 16.5 | 63.1542 | 6.0796
13.6084 | 17 | 63.3994 | 6.2221
13.7688 | 17.5 | 63.6675 | 6.3582
13.8578 | 18 | 63.815 | 6.5133
14.0372 | 18.5 | 64.1094 | 6.6398
14.1493 | 19 | 64.2914 | 6.7846
14.2994 | 19.5 | 64.5329 | 6.9161
14.4492 | 20 | 64.7714 | 7.0457
14.7235 | 21 | 65.2018 | 7.3076
14.9235 | 22 | 65.5105 | 7.5877
15.1589 | 23 | 65.8687 | 7.8502
15.3907 | 24 | 66.216 | 8.1081
15.5651 | 25 | 66.4738 | 8.3816
15.8793 | 26 | 66.9312 | 8.5979
16.0865 | 27 | 67.2279 | 8.8485
16.2894 | 28 | 67.5148 | 9.0959
16.4879 | 29 | 67.7919 | 9.3403
16.6816 | 30 | 68.0593 | 9.5822
16.8721 | 31 | 68.319 | 9.8211
17.0592 | 32 | 68.5714 | 10.0571
17.2438 | 33 | 68.8177 | 10.2902
17.4243 | 34 | 69.0561 | 10.5209
17.5939 | 35 | 69.2777 | 10.7528
17.785 | 36 | 69.525 | 10.971
17.952 | 37 | 69.7389 | 11.1966
18.1212 | 38 | 69.9536 | 11.4176
18.2867 | 39 | 70.1616 | 11.637
18.4526 | 40 | 70.3683 | 11.8527
18.6207 | 41 | 70.5758 | 12.0639
18.7848 | 42 | 70.7766 | 12.2738
18.9492 | 43 | 70.976 | 12.4803
19.092 | 44 | 71.1479 | 12.6949
19.2495 | 45 | 71.3359 | 12.8989
19.4125 | 46 | 71.5288 | 13.0968
19.5623 | 47 | 71.7047 | 13.2988
19.7106 | 48 | 71.8775 | 13.4988
19.8611 | 49 | 72.0516 | 13.6947
20.0134 | 50 | 72.2265 | 13.8868
20.1515 | 51 | 72.3838 | 14.0843
20.3035 | 52 | 72.5558 | 14.271
20.443 | 53 | 72.7125 | 14.4624
20.5657 | 54 | 72.8494 | 14.6613
20.7378 | 55 | 73.0402 | 14.8279
20.8513 | 56 | 73.1651 | 15.0275
20.996 | 57 | 73.3233 | 15.2057
21.1047 | 58 | 73.4415 | 15.4039
21.2856 | 59 | 73.6369 | 15.5542
21.382 | 60 | 73.7402 | 15.7559
21.4801 | 61 | 73.845 | 15.9545
21.6661 | 62 | 74.0424 | 16.0937
21.7701 | 63 | 74.1519 | 16.2843
21.8342 | 64 | 74.2192 | 16.4997
22.0547 | 65 | 74.4491 | 16.6081
22.1667 | 66 | 74.5651 | 16.787
22.2627 | 67 | 74.664 | 16.9752
22.3312 | 68 | 74.7342 | 17.1807
22.4348 | 69 | 74.8402 | 17.3603
22.5414 | 70 | 74.9486 | 17.5359
22.6238 | 71 | 75.0322 | 17.7272
22.8593 | 72 | 75.2691 | 17.8062
22.8743 | 73 | 75.2841 | 18.0426
22.8596 | 74 | 75.2695 | 18.3006
22.9655 | 75 | 75.3752 | 18.4686
23.2324 | 76 | 75.6397 | 18.5138
23.2339 | 77 | 75.6411 | 18.7563
23.38 | 78 | 75.7846 | 18.888
23.3284 | 79 | 75.734 | 19.1701
23.4419 | 80 | 75.8451 | 19.3239
23.644 | 81 | 76.0416 | 19.4063
23.6689 | 82 | 76.0657 | 19.6261
23.7407 | 83 | 76.135 | 19.808
23.7301 | 84 | 76.1247 | 20.0552
23.9952 | 85 | 76.379 | 20.0778
23.8282 | 86 | 76.2192 | 20.4515
24.1403 | 87 | 76.5169 | 20.4303
24.2256 | 88 | 76.5977 | 20.5941
24.2959 | 89 | 76.664 | 20.7691
24.2662 | 90 | 76.636 | 21.0276
24.3801 | 91 | 76.7432 | 21.1637
24.5259 | 92 | 76.8796 | 21.2708
24.4958 | 93 | 76.8516 | 21.5281
24.3647 | 94 | 76.7287 | 21.875
24.5405 | 95 | 76.8933 | 21.9514
24.8954 | 96 | 77.2218 | 21.8671
24.8694 | 97 | 77.1979 | 22.118
24.6888 | 98 | 77.0312 | 22.5095
24.8445 | 99 | 77.175 | 22.5968
25.0869 | 100 | 77.3972 | 22.6028
25.2931 | 105 | 77.5845 | 23.5363
25.6952 | 110 | 77.9455 | 24.26
25.782 | 115 | 78.0226 | 25.274
26.0577 | 120 | 78.2661 | 26.0807
26.484 | 125 | 78.6374 | 26.7032
26.688 | 130 | 78.813 | 27.5431
26.8129 | 135 | 78.9199 | 28.4582
26.9996 | 140 | 79.0787 | 29.2898
27.4061 | 145 | 79.4207 | 29.84
27.472 | 150 | 79.4757 | 30.7865
28.4678 | 155 | 80.2905 | 30.5498
27.4519 | 160 | 79.4589 | 32.8658
27.5629 | 165 | 79.5512 | 33.7405
28.2308 | 170 | 80.0992 | 33.8314
28.0298 | 175 | 79.9356 | 35.1126
28.214 | 180 | 80.0855 | 35.846
28.0151 | 185 | 79.9237 | 37.1412
28.4897 | 190 | 80.3081 | 37.4146
28.8347 | 195 | 80.5836 | 37.8621
28.5568 | 200 | 80.3619 | 39.2762
29.0804 | 205 | 80.7778 | 39.4056
29.1885 | 210 | 80.8626 | 40.1885
29.6569 | 215 | 81.2269 | 40.3621
29.3822 | 220 | 81.014 | 41.7692
29.6953 | 225 | 81.2566 | 42.1727
29.6949 | 230 | 81.2563 | 43.1105
29.9451 | 235 | 81.4483 | 43.5966
29.9557 | 240 | 81.4564 | 44.5047
29.7766 | 245 | 81.3192 | 45.7681
29.9022 | 250 | 81.4155 | 46.4613
29.6913 | 255 | 81.2535 | 47.8036
30.405 | 260 | 81.7971 | 47.3275
30.8251 | 265 | 82.1111 | 47.4056
30.6581 | 270 | 81.9868 | 48.6357
30.6321 | 275 | 81.9674 | 49.5897
30.6317 | 280 | 81.9671 | 50.4921
30.8145 | 285 | 82.1032 | 51.0059
31.0734 | 290 | 82.2947 | 51.3454
31.1027 | 295 | 82.3163 | 52.167
31.1733 | 300 | 82.3682 | 52.8954
31.4865 | 310 | 82.5969 | 53.9496
31.7714 | 320 | 82.8031 | 55.0302
31.8926 | 330 | 82.8902 | 56.4623
32.6133 | 340 | 83.4016 | 56.4345
32.3108 | 350 | 83.1884 | 58.8407
33.0784 | 360 | 83.7257 | 58.5876
32.751 | 370 | 83.498 | 61.0573
33.0234 | 380 | 83.6876 | 61.987
33.1516 | 390 | 83.7763 | 63.2725
33.3864 | 400 | 83.9377 | 64.249
33.624 | 410 | 84.1001 | 65.1896
33.6639 | 420 | 84.1272 | 66.6656
33.913 | 430 | 84.2959 | 67.5276
33.9818 | 440 | 84.3423 | 68.8938
34.1777 | 450 | 84.4739 | 69.8675
34.4253 | 460 | 84.639 | 70.6604
34.5374 | 470 | 84.7134 | 71.8469
34.7054 | 480 | 84.8245 | 72.8423
34.838 | 490 | 84.9118 | 73.9323
35.2815 | 500 | 85.2013 | 73.9937
35.1332 | 520 | 85.1049 | 77.4548
35.532 | 540 | 85.3632 | 79.0389
36.0707 | 560 | 85.7075 | 80.0379
36.1443 | 580 | 85.7542 | 82.6258
36.4226 | 600 | 85.9297 | 84.4217
36.6892 | 620 | 86.0966 | 86.201
36.947 | 640 | 86.2569 | 87.9561
37.2799 | 660 | 86.4621 | 89.3499
37.5787 | 680 | 86.6448 | 90.8151
37.8204 | 700 | 86.7916 | 92.4591
38.4818 | 720 | 87.1883 | 92.2442
38.5803 | 740 | 87.2468 | 94.3736
38.3276 | 760 | 87.0964 | 98.0671
39.5929 | 780 | 87.8397 | 94.8504
39.657 | 800 | 87.8767 | 96.9863
40.0162 | 820 | 88.0831 | 97.7186
39.7906 | 840 | 87.9537 | 101.1888
40.6238 | 860 | 88.4279 | 99.5198
41.0079 | 880 | 88.6433 | 99.9387
41.4119 | 900 | 88.8676 | 100.1912
41.8799 | 920 | 89.1248 | 100.0516
42.1835 | 940 | 89.2901 | 100.6729
41.941 | 960 | 89.1582 | 104.0811
42.7603 | 980 | 89.6009 | 101.9111
43.025 | 1000 | 89.7422 | 102.5782
43.7849 | 1050 | 90.1428 | 103.5002
44.239 | 1100 | 90.379 | 105.8315
45.3962 | 1150 | 90.9699 | 103.8462
45.8405 | 1200 | 91.1927 | 105.687
46.3967 | 1250 | 91.4688 | 106.6402
47.2383 | 1300 | 91.8802 | 105.558
47.8605 | 1350 | 92.1796 | 105.5752
48.4329 | 1400 | 92.4517 | 105.6764
49.0408 | 1450 | 92.7371 | 105.3114
49.5469 | 1500 | 92.9721 | 105.4184
53.9568 | 2000 | 94.9234 | 101.5324
59.9996 | 3000 | 97.3527 | 79.4179
64.9757 | 4000 | 99.1761 | 32.9561
68.8085 | 5000 | 100.4877 | 5024.3858
72.1791 | 6000 | 101.5822 | 6094.9292
72.6888 | 6210 | 101.7432 | 6318.2534
77.9542 | 8000 | 103.3436 | 8267.4918
82.8611 | 10000 | 104.7406 | 10474.0633
87.5944 | 12000 | 106.0119 | 12721.4304
89.3463 | 12800 | 106.4651 | 13627.5353
91.8724 | 14000 | 107.1032 | 14994.4449
95.7458 | 16000 | 108.0482 | 17287.7188
99.3713 | 18000 | 108.8988 | 19601.7833
102.8276 | 20000 | 109.6812 | 21936.2489
106.1111 | 22000 | 110.4006 | 24288.1262
109.2389 | 24000 | 111.0654 | 26655.6954
111.9335 | 25700 | 111.623 | 28687.1236
118.0613 | 30000 | 112.8428 | 33852.8397
124.6011 | 35000 | 114.0766 | 39926.8127
130.1517 | 39400 | 115.074 | 45339.1644
136.845 | 45000 | 116.2217 | 52299.7435
142.3674 | 50000 | 117.127 | 58563.5209
148.9149 | 56300 | 118.156 | 66521.8503
152.5902 | 60000 | 118.714 | 71228.3971
157.3773 | 65000 | 119.4209 | 77623.5973
162.0293 | 70000 | 120.0876 | 84061.3063
168.8503 | 77700 | 121.0313 | 94041.2869
179.4261 | 90000 | 122.4215 | 110179.3878
187.2968 | 100000 | 123.404 | 123404.0084
194.8923 | 110000 | 124.3138 | 136745.1314

## Rating Curve Calibration 2
DateTime | PT Stage (in) | Manually Measured Flow (cfs) | Sutron Flow (cfs) | PT Stage Corresponding to Manually Measured Flow (in) | Absolute Error in Level (in) | Revised Flow (cfs) | Absolute Error (cfs)
2020-03-26 15:00 | 2.79 | 0.988 | 1.8441 | 2.0385 | 0.75147 | 2.6156 | 1.6276
2020-03-31 10:20 | 2.67 | 1.8647 | 1.6919 | 2.8062 | -0.13621 | 2.5031 | 0.63842
2020-04-07 09:10 | 5.73 | 4.4089 | 7.5035 | 4.3472 | 1.3828 | 5.3719 | 0.96302
2020-04-08 15:15 | 10.03 | 12.8269 | 8.2937 | 12.1842 | -2.1542 | 9.4031 | 3.4238
2020-04-09 11:20 | 38.67 | 87.1792 | 765.4118 | 24.1555 | 14.5145 | 86.1963 | 0.98291
2020-04-09 12:05 | 39.79 | 97.3894 | 807.4052 | 24.7991 | 14.9909 | 99.2259 | 1.8365
2020-03-12 14:00 | 0.04 | 0 | 0.013784 | 0 | 0.04 | 0.0375 | 0.0375
2020-06-11 15:23 | 1.66 | 0.003984 | 0.67363 | 0.011561 | 1.6484 | 1.5562 | 1.5523
2020-06-11 15:37 | 1.57 | 0.0030412 | 0.59889 | 0.0088251 | 1.5612 | 1.4719 | 1.4688
2020-06-11 15:46 | 1.55 | 0.0041969 | 0.58228 | 0.012179 | 1.5378 | 1.4531 | 1.4489
2020-06-16 14:45 | 1.4 | 0.0039077 | 0.48245 | 0.01134 | 1.3887 | 1.3125 | 1.3086
2020-06-16 14:50 | 1.39 | 0.0035294 | 0.479 | 0.010242 | 1.3798 | 1.3031 | 1.2996
HEC-RAS rating curve |  |  |  | Mean Absolute Error (in) | 4.1985 | Mean Absolute Error (cfs) | 1.3585
Stage (in) | Q Total (cfs) | Revised Stage (in) | Revised Flow (cfs) | Horizontal Shift to Rating Curve (in) | 16
0 | 0 | 0 | 0 | Vertical Shift to Rating Curve (cfs) | 15
1.4509 | 0.5 | 16 | 15 | Difference between maximum absolute errors (cfs) | 0.66461
2.053 | 1 | 17.4509 | 15.5
2.5188 | 1.5 | 18.053 | 16
2.9128 | 2 | 18.5188 | 16.5
3.2607 | 2.5 | 18.9128 | 17
3.576 | 3 | 19.2607 | 17.5
3.8672 | 3.5 | 19.576 | 18
4.1382 | 4 | 19.8672 | 18.5
4.3938 | 4.5 | 20.1382 | 19
4.6355 | 5 | 20.3938 | 19.5
4.8662 | 5.5 | 20.6355 | 20
5.087 | 6 | 20.8662 | 20.5
5.2991 | 6.5 | 21.087 | 21
5.5038 | 7 | 21.2991 | 21.5
5.7012 | 7.5 | 21.5038 | 22
9.8452 | 8 | 21.7012 | 22.5
10.1598 | 8.5 | 25.8452 | 23
10.4498 | 9 | 26.1598 | 23.5
10.7153 | 9.5 | 26.4498 | 24
10.9629 | 10 | 26.7153 | 24.5
11.2101 | 10.5 | 26.9629 | 25
11.4382 | 11 | 27.2101 | 25.5
11.6503 | 11.5 | 27.4382 | 26
11.859 | 12 | 27.6503 | 26.5
12.0582 | 12.5 | 27.859 | 27
12.2509 | 13 | 28.0582 | 27.5
12.4365 | 13.5 | 28.2509 | 28
12.6167 | 14 | 28.4365 | 28.5
12.8071 | 14.5 | 28.6167 | 29
12.9562 | 15 | 28.8071 | 29.5
13.1191 | 15.5 | 28.9562 | 30
13.277 | 16 | 29.1191 | 30.5
13.4634 | 16.5 | 29.277 | 31
13.6084 | 17 | 29.4634 | 31.5
13.7688 | 17.5 | 29.6084 | 32
13.8578 | 18 | 29.7688 | 32.5
14.0372 | 18.5 | 29.8578 | 33
14.1493 | 19 | 30.0372 | 33.5
14.2994 | 19.5 | 30.1493 | 34
14.4492 | 20 | 30.2994 | 34.5
14.7235 | 21 | 30.4492 | 35
14.9235 | 22 | 30.7235 | 36
15.1589 | 23 | 30.9235 | 37
15.3907 | 24 | 31.1589 | 38
15.5651 | 25 | 31.3907 | 39
15.8793 | 26 | 31.5651 | 40
16.0865 | 27 | 31.8793 | 41
16.2894 | 28 | 32.0865 | 42
16.4879 | 29 | 32.2894 | 43
16.6816 | 30 | 32.4879 | 44
16.8721 | 31 | 32.6816 | 45
17.0592 | 32 | 32.8721 | 46
17.2438 | 33 | 33.0592 | 47
17.4243 | 34 | 33.2438 | 48
17.5939 | 35 | 33.4243 | 49
17.785 | 36 | 33.5939 | 50
17.952 | 37 | 33.785 | 51
18.1212 | 38 | 33.952 | 52
18.2867 | 39 | 34.1212 | 53
18.4526 | 40 | 34.2867 | 54
18.6207 | 41 | 34.4526 | 55
18.7848 | 42 | 34.6207 | 56
18.9492 | 43 | 34.7848 | 57
19.092 | 44 | 34.9492 | 58
19.2495 | 45 | 35.092 | 59
19.4125 | 46 | 35.2495 | 60
19.5623 | 47 | 35.4125 | 61
19.7106 | 48 | 35.5623 | 62
19.8611 | 49 | 35.7106 | 63
20.0134 | 50 | 35.8611 | 64
20.1515 | 51 | 36.0134 | 65
20.3035 | 52 | 36.1515 | 66
20.443 | 53 | 36.3035 | 67
20.5657 | 54 | 36.443 | 68
20.7378 | 55 | 36.5657 | 69
20.8513 | 56 | 36.7378 | 70
20.996 | 57 | 36.8513 | 71
21.1047 | 58 | 36.996 | 72
21.2856 | 59 | 37.1047 | 73
21.382 | 60 | 37.2856 | 74
21.4801 | 61 | 37.382 | 75
21.6661 | 62 | 37.4801 | 76
21.7701 | 63 | 37.6661 | 77
21.8342 | 64 | 37.7701 | 78
22.0547 | 65 | 37.8342 | 79
22.1667 | 66 | 38.0547 | 80
22.2627 | 67 | 38.1667 | 81
22.3312 | 68 | 38.2627 | 82
22.4348 | 69 | 38.3312 | 83
22.5414 | 70 | 38.4348 | 84
22.6238 | 71 | 38.5414 | 85
22.8593 | 72 | 38.6238 | 86
22.8743 | 73 | 38.8593 | 87
22.8596 | 74 | 38.8743 | 88
22.9655 | 75 | 38.8596 | 89
23.2324 | 76 | 38.9655 | 90
23.2339 | 77 | 39.2324 | 91
23.38 | 78 | 39.2339 | 92
23.3284 | 79 | 39.38 | 93
23.4419 | 80 | 39.3284 | 94
23.644 | 81 | 39.4419 | 95
23.6689 | 82 | 39.644 | 96
23.7407 | 83 | 39.6689 | 97
23.7301 | 84 | 39.7407 | 98
23.9952 | 85 | 39.7301 | 99
23.8282 | 86 | 39.9952 | 100
24.1403 | 87 | 39.8282 | 101
24.2256 | 88 | 40.1403 | 102
24.2959 | 89 | 40.2256 | 103
24.2662 | 90 | 40.2959 | 104
24.3801 | 91 | 40.2662 | 105
24.5259 | 92 | 40.3801 | 106
24.4958 | 93 | 40.5259 | 107
24.3647 | 94 | 40.4958 | 108
24.5405 | 95 | 40.3647 | 109
24.8954 | 96 | 40.5405 | 110
24.8694 | 97 | 40.8954 | 111
24.6888 | 98 | 40.8694 | 112
24.8445 | 99 | 40.6888 | 113
25.0869 | 100 | 40.8445 | 114
25.2931 | 105 | 41.0869 | 115
25.6952 | 110 | 41.2931 | 120
25.782 | 115 | 41.6952 | 125
26.0577 | 120 | 41.782 | 130
26.484 | 125 | 42.0577 | 135
26.688 | 130 | 42.484 | 140
26.8129 | 135 | 42.688 | 145
26.9996 | 140 | 42.8129 | 150
27.4061 | 145 | 42.9996 | 155
27.472 | 150 | 43.4061 | 160
28.4678 | 155 | 43.472 | 165
27.4519 | 160 | 44.4678 | 170
27.5629 | 165 | 43.4519 | 175
28.2308 | 170 | 43.5629 | 180
28.0298 | 175 | 44.2308 | 185
28.214 | 180 | 44.0298 | 190
28.0151 | 185 | 44.214 | 195
28.4897 | 190 | 44.0151 | 200
28.8347 | 195 | 44.4897 | 205
28.5568 | 200 | 44.8347 | 210
29.0804 | 205 | 44.5568 | 215
29.1885 | 210 | 45.0804 | 220
29.6569 | 215 | 45.1885 | 225
29.3822 | 220 | 45.6569 | 230
29.6953 | 225 | 45.3822 | 235
29.6949 | 230 | 45.6953 | 240
29.9451 | 235 | 45.6949 | 245
29.9557 | 240 | 45.9451 | 250
29.7766 | 245 | 45.9557 | 255
29.9022 | 250 | 45.7766 | 260
29.6913 | 255 | 45.9022 | 265
30.405 | 260 | 45.6913 | 270
30.8251 | 265 | 46.405 | 275
30.6581 | 270 | 46.8251 | 280
30.6321 | 275 | 46.6581 | 285
30.6317 | 280 | 46.6321 | 290
30.8145 | 285 | 46.6317 | 295
31.0734 | 290 | 46.8145 | 300
31.1027 | 295 | 47.0734 | 305
31.1733 | 300 | 47.1027 | 310
31.4865 | 310 | 47.1733 | 315
31.7714 | 320 | 47.4865 | 325
31.8926 | 330 | 47.7714 | 335
32.6133 | 340 | 47.8926 | 345
32.3108 | 350 | 48.6133 | 355
33.0784 | 360 | 48.3108 | 365
32.751 | 370 | 49.0784 | 375
33.0234 | 380 | 48.751 | 385
33.1516 | 390 | 49.0234 | 395
33.3864 | 400 | 49.1516 | 405
33.624 | 410 | 49.3864 | 415
33.6639 | 420 | 49.624 | 425
33.913 | 430 | 49.6639 | 435
33.9818 | 440 | 49.913 | 445
34.1777 | 450 | 49.9818 | 455
34.4253 | 460 | 50.1777 | 465
34.5374 | 470 | 50.4253 | 475
34.7054 | 480 | 50.5374 | 485
34.838 | 490 | 50.7054 | 495
35.2815 | 500 | 50.838 | 505
35.1332 | 520 | 51.2815 | 515
35.532 | 540 | 51.1332 | 535
36.0707 | 560 | 51.532 | 555
36.1443 | 580 | 52.0707 | 575
36.4226 | 600 | 52.1443 | 595
36.6892 | 620 | 52.4226 | 615
36.947 | 640 | 52.6892 | 635
37.2799 | 660 | 52.947 | 655
37.5787 | 680 | 53.2799 | 675
37.8204 | 700 | 53.5787 | 695
38.4818 | 720 | 53.8204 | 715
38.5803 | 740 | 54.4818 | 735
38.3276 | 760 | 54.5803 | 755
39.5929 | 780 | 54.3276 | 775
39.657 | 800 | 55.5929 | 795
40.0162 | 820 | 55.657 | 815
39.7906 | 840 | 56.0162 | 835
40.6238 | 860 | 55.7906 | 855
41.0079 | 880 | 56.6238 | 875
41.4119 | 900 | 57.0079 | 895
41.8799 | 920 | 57.4119 | 915
42.1835 | 940 | 57.8799 | 935
41.941 | 960 | 58.1835 | 955
42.7603 | 980 | 57.941 | 975
43.025 | 1000 | 58.7603 | 995
43.7849 | 1050 | 59.025 | 1015
44.239 | 1100 | 59.7849 | 1065
45.3962 | 1150 | 60.239 | 1115
45.8405 | 1200 | 61.3962 | 1165
46.3967 | 1250 | 61.8405 | 1215
47.2383 | 1300 | 62.3967 | 1265
47.8605 | 1350 | 63.2383 | 1315
48.4329 | 1400 | 63.8605 | 1365
49.0408 | 1450 | 64.4329 | 1415
49.5469 | 1500 | 65.0408 | 1465
53.9568 | 2000 | 65.5469 | 1515
59.9996 | 3000 | 69.9568 | 2015
64.9757 | 4000 | 75.9996 | 3015
68.8085 | 5000 | 80.9757 | 4015
72.1791 | 6000 | 84.8085 | 5015
72.6888 | 6210 | 88.1791 | 6015
77.9542 | 8000 | 88.6888 | 6225
82.8611 | 10000 | 93.9542 | 8015
87.5944 | 12000 | 98.8611 | 10015
89.3463 | 12800 | 103.5944 | 12015
91.8724 | 14000 | 105.3463 | 12815
95.7458 | 16000 | 107.8724 | 14015
99.3713 | 18000 | 111.7458 | 16015
102.8276 | 20000 | 115.3713 | 18015
106.1111 | 22000 | 118.8276 | 20015
109.2389 | 24000 | 122.1111 | 22015
111.9335 | 25700 | 125.2389 | 24015
118.0613 | 30000 | 127.9335 | 25715
124.6011 | 35000 | 134.0613 | 30015
130.1517 | 39400 | 140.6011 | 35015
136.845 | 45000 | 146.1517 | 39415
142.3674 | 50000 | 152.845 | 45015
148.9149 | 56300 | 158.3674 | 50015
152.5902 | 60000 | 164.9149 | 56315
157.3773 | 65000 | 168.5902 | 60015
162.0293 | 70000 | 173.3773 | 65015
168.8503 | 77700 | 178.0293 | 70015
179.4261 | 90000 | 184.8503 | 77715
187.2968 | 100000 | 195.4261 | 90015
194.8923 | 110000 | 203.2968 | 100015
|  | 210.8923 | 110015

## 4x24in CMP Rating Curve
MANNING'S EQUATION FOR PIPE FLOW
INPUT
[7 empty] | 24 | Pipe diameter (in)
| Mannings Formula | [5 empty] | 0.021 | mannings coeff
| Q=(1.486/n)ARh2/3S1/2 | [5 empty] | 0.005 | slope in/in
R=A/P
A=cross sectional area
|  | P=wetted perimeter |  |  |  | V=(1.49/n)Rh2/3S1/2
|  | S=slope of channel |  |  |  | Q=V x A
n=Manning's roughness coefficient
level (in) | angle | Area,ft2 | Wetted Perimeter, ft | Hydraulic Radius, ft | velocity ft/s | flow, cfs | Estimated PT Stage
0.1 | 14.804 | 0.0014326 | 0.25838 | 0.0055447 | 0.15672 | 0.00089811 | 0.1
0.2 | 20.9506 | 0.004047 | 0.36566 | 0.011068 | 0.24846 | 0.0040222 | 0.2
0.3 | 25.6771 | 0.0074255 | 0.44815 | 0.016569 | 0.32516 | 0.0096579 | 0.3
0.4 | 29.6702 | 0.011418 | 0.51784 | 0.022049 | 0.39339 | 0.017967 | 0.4
0.5 | 33.1957 | 0.015937 | 0.57937 | 0.027507 | 0.45589 | 0.029062 | 0.5
0.6 | 36.3897 | 0.020923 | 0.63512 | 0.032943 | 0.51414 | 0.043029 | 0.6
0.7 | 39.3333 | 0.026332 | 0.6865 | 0.038358 | 0.56904 | 0.059937 | 0.7
0.8 | 42.0789 | 0.032131 | 0.73442 | 0.043751 | 0.6212 | 0.079839 | 0.8
0.9 | 44.6633 | 0.038291 | 0.77952 | 0.049121 | 0.67105 | 0.10278 | 0.9
1 | 47.1129 | 0.04479 | 0.82228 | 0.05447 | 0.71892 | 0.1288 | 1
1.1 | 49.4479 | 0.051607 | 0.86303 | 0.059797 | 0.76507 | 0.15793 | 1.1
1.2 | 51.6839 | 0.058726 | 0.90205 | 0.065102 | 0.80967 | 0.1902 | 1.2
1.3 | 53.8331 | 0.066132 | 0.93957 | 0.070386 | 0.85291 | 0.22562 | 1.3
1.4 | 55.9058 | 0.073811 | 0.97574 | 0.075647 | 0.8949 | 0.26422 | 1.4
1.5 | 57.91 | 0.081753 | 1.0107 | 0.080886 | 0.93576 | 0.306 | 1.5
1.6 | 59.8529 | 0.089946 | 1.0446 | 0.086103 | 0.97557 | 0.35099 | 1.6
1.7 | 61.7401 | 0.098379 | 1.0776 | 0.091298 | 1.0144 | 0.3992 | 1.7
1.8 | 63.5767 | 0.10705 | 1.1096 | 0.09647 | 1.0524 | 0.45062 | 1.8
1.9 | 65.3669 | 0.11594 | 1.1409 | 0.10162 | 1.0895 | 0.50527 | 1.9
2 | 67.1146 | 0.12504 | 1.1714 | 0.10675 | 1.1259 | 0.56314 | 2
2.1 | 68.823 | 0.13436 | 1.2012 | 0.11186 | 1.1615 | 0.62424 | 2.1
2.2 | 70.495 | 0.14388 | 1.2304 | 0.11694 | 1.1964 | 0.68857 | 2.2
2.3 | 72.1332 | 0.15359 | 1.259 | 0.122 | 1.2307 | 0.75613 | 2.3
2.4 | 73.7398 | 0.1635 | 1.287 | 0.12704 | 1.2644 | 0.82692 | 2.4
2.5 | 75.3169 | 0.17359 | 1.3145 | 0.13206 | 1.2975 | 0.90092 | 2.5
2.6 | 76.8664 | 0.18386 | 1.3416 | 0.13705 | 1.33 | 0.97814 | 2.6
2.7 | 78.3899 | 0.19431 | 1.3682 | 0.14202 | 1.362 | 1.0586 | 2.7
2.8 | 79.889 | 0.20493 | 1.3943 | 0.14697 | 1.3934 | 1.1422 | 2.8
2.9 | 81.365 | 0.21571 | 1.4201 | 0.1519 | 1.4244 | 1.229 | 2.9
3 | 82.8192 | 0.22666 | 1.4455 | 0.1568 | 1.4549 | 1.319 | 3
3.1 | 84.2528 | 0.23776 | 1.4705 | 0.16169 | 1.4849 | 1.4122 | 3.1
3.2 | 85.6669 | 0.24901 | 1.4952 | 0.16655 | 1.5145 | 1.5086 | 3.2
3.3 | 87.0623 | 0.26042 | 1.5195 | 0.17138 | 1.5437 | 1.6081 | 3.3
3.4 | 88.4401 | 0.27197 | 1.5436 | 0.1762 | 1.5725 | 1.7107 | 3.4
3.5 | 89.8011 | 0.28367 | 1.5673 | 0.18099 | 1.6009 | 1.8165 | 3.5
3.6 | 91.146 | 0.2955 | 1.5908 | 0.18576 | 1.6289 | 1.9253 | 3.6
3.7 | 92.4756 | 0.30747 | 1.614 | 0.1905 | 1.6565 | 2.0373 | 3.7
3.8 | 93.7907 | 0.31957 | 1.637 | 0.19522 | 1.6838 | 2.1523 | 3.8
3.9 | 95.0917 | 0.3318 | 1.6597 | 0.19992 | 1.7107 | 2.2705 | 3.9
4 | 96.3794 | 0.34416 | 1.6821 | 0.2046 | 1.7373 | 2.3916 | 4
4.1 | 97.6542 | 0.35665 | 1.7044 | 0.20925 | 1.7635 | 2.5158 | 4.1
4.2 | 98.9168 | 0.36925 | 1.7264 | 0.21388 | 1.7894 | 2.643 | 4.2
4.3 | 100.1676 | 0.38198 | 1.7483 | 0.21849 | 1.815 | 2.7732 | 4.3
4.4 | 101.407 | 0.39482 | 1.7699 | 0.22308 | 1.8403 | 2.9064 | 4.4
4.5 | 102.6356 | 0.40777 | 1.7913 | 0.22764 | 1.8653 | 3.0426 | 4.5
4.6 | 103.8538 | 0.42084 | 1.8126 | 0.23218 | 1.8901 | 3.1816 | 4.6
4.7 | 105.0618 | 0.43401 | 1.8337 | 0.23669 | 1.9145 | 3.3237 | 4.7
4.8 | 106.2602 | 0.4473 | 1.8546 | 0.24118 | 1.9386 | 3.4686 | 4.8
4.9 | 107.4493 | 0.46068 | 1.8753 | 0.24565 | 1.9625 | 3.6163 | 4.9
5 | 108.6293 | 0.47417 | 1.8959 | 0.2501 | 1.9861 | 3.767 | 5
5.1 | 109.8007 | 0.48775 | 1.9164 | 0.25452 | 2.0094 | 3.9205 | 5.1
5.2 | 110.9638 | 0.50144 | 1.9367 | 0.25892 | 2.0325 | 4.0768 | 5.2
5.3 | 112.1188 | 0.51522 | 1.9568 | 0.26329 | 2.0554 | 4.2358 | 5.3
5.4 | 113.266 | 0.52909 | 1.9769 | 0.26764 | 2.0779 | 4.3977 | 5.4
5.5 | 114.4057 | 0.54306 | 1.9968 | 0.27197 | 2.1003 | 4.5623 | 5.5
5.6 | 115.5381 | 0.55711 | 2.0165 | 0.27627 | 2.1224 | 4.7296 | 5.6
5.7 | 116.6635 | 0.57125 | 2.0362 | 0.28055 | 2.1443 | 4.8996 | 5.7
5.8 | 117.7822 | 0.58548 | 2.0557 | 0.28481 | 2.1659 | 5.0723 | 5.8
5.9 | 118.8942 | 0.59979 | 2.0751 | 0.28904 | 2.1873 | 5.2477 | 5.9
6 | 120 | 0.61418 | 2.0944 | 0.29325 | 2.2085 | 5.4257 | 6
6.1 | 121.0996 | 0.62866 | 2.1136 | 0.29744 | 2.2294 | 5.6062 | 6.1
6.2 | 122.1933 | 0.64321 | 2.1327 | 0.3016 | 2.2502 | 5.7894 | 6.2
6.3 | 123.2813 | 0.65784 | 2.1517 | 0.30573 | 2.2707 | 5.9751 | 6.3
6.4 | 124.3637 | 0.67254 | 2.1706 | 0.30985 | 2.291 | 6.1633 | 6.4
6.5 | 125.4408 | 0.68732 | 2.1894 | 0.31394 | 2.3112 | 6.354 | 6.5
6.6 | 126.5126 | 0.70217 | 2.2081 | 0.318 | 2.3311 | 6.5472 | 6.6
6.7 | 127.5795 | 0.71709 | 2.2267 | 0.32204 | 2.3508 | 6.7428 | 6.7
6.8 | 128.6414 | 0.73207 | 2.2452 | 0.32606 | 2.3703 | 6.9409 | 6.8
6.9 | 129.6987 | 0.74713 | 2.2637 | 0.33005 | 2.3896 | 7.1413 | 6.9
7 | 130.7514 | 0.76225 | 2.282 | 0.33402 | 2.4087 | 7.3441 | 7
7.1 | 131.7996 | 0.77743 | 2.3003 | 0.33796 | 2.4276 | 7.5492 | 7.1
7.2 | 132.8436 | 0.79267 | 2.3186 | 0.34188 | 2.4464 | 7.7566 | 7.2
7.3 | 133.8835 | 0.80798 | 2.3367 | 0.34578 | 2.4649 | 7.9663 | 7.3
7.4 | 134.9194 | 0.82334 | 2.3548 | 0.34965 | 2.4833 | 8.1783 | 7.4
7.5 | 135.9514 | 0.83877 | 2.3728 | 0.35349 | 2.5014 | 8.3925 | 7.5
7.6 | 136.9796 | 0.85424 | 2.3907 | 0.35731 | 2.5194 | 8.6088 | 7.6
7.7 | 138.0042 | 0.86978 | 2.4086 | 0.36111 | 2.5372 | 8.8273 | 7.7
7.8 | 139.0254 | 0.88536 | 2.4265 | 0.36488 | 2.5549 | 9.048 | 7.8
7.9 | 140.0431 | 0.901 | 2.4442 | 0.36863 | 2.5723 | 9.2707 | 7.9
8 | 141.0576 | 0.91669 | 2.4619 | 0.37235 | 2.5896 | 9.4955 | 8
8.1 | 142.0689 | 0.93243 | 2.4796 | 0.37604 | 2.6067 | 9.7224 | 8.1
8.2 | 143.0771 | 0.94821 | 2.4972 | 0.37972 | 2.6237 | 9.9512 | 8.2
8.3 | 144.0824 | 0.96404 | 2.5147 | 0.38336 | 2.6404 | 10.182 | 8.3
8.4 | 145.0848 | 0.97992 | 2.5322 | 0.38698 | 2.657 | 10.4148 | 8.4
8.5 | 146.0845 | 0.99584 | 2.5497 | 0.39058 | 2.6735 | 10.6495 | 8.5
8.6 | 147.0815 | 1.0118 | 2.5671 | 0.39415 | 2.6898 | 10.886 | 8.6
8.7 | 148.076 | 1.0278 | 2.5844 | 0.3977 | 2.7059 | 11.1244 | 8.7
8.8 | 149.068 | 1.0439 | 2.6017 | 0.40122 | 2.7218 | 11.3646 | 8.8
8.9 | 150.0576 | 1.0599 | 2.619 | 0.40471 | 2.7376 | 11.6066 | 8.9
9 | 151.045 | 1.0761 | 2.6362 | 0.40818 | 2.7532 | 11.8504 | 9
9.1 | 152.0301 | 1.0922 | 2.6534 | 0.41162 | 2.7687 | 12.0959 | 9.1
9.2 | 153.0132 | 1.1084 | 2.6706 | 0.41504 | 2.784 | 12.343 | 9.2
9.3 | 153.9942 | 1.1246 | 2.6877 | 0.41843 | 2.7991 | 12.5918 | 9.3
9.4 | 154.9733 | 1.1409 | 2.7048 | 0.4218 | 2.8141 | 12.8422 | 9.4
9.5 | 155.9506 | 1.1572 | 2.7219 | 0.42514 | 2.829 | 13.0942 | 9.5
9.6 | 156.9261 | 1.1735 | 2.7389 | 0.42845 | 2.8436 | 13.3478 | 9.6
9.7 | 157.8999 | 1.1898 | 2.7559 | 0.43174 | 2.8582 | 13.6029 | 9.7
9.8 | 158.872 | 1.2062 | 2.7728 | 0.435 | 2.8725 | 13.8594 | 9.8
9.9 | 159.8427 | 1.2226 | 2.7898 | 0.43824 | 2.8868 | 14.1174 | 9.9
10 | 160.8119 | 1.239 | 2.8067 | 0.44145 | 2.9009 | 14.3768 | 10
10.1 | 161.7797 | 1.2555 | 2.8236 | 0.44463 | 2.9148 | 14.6376 | 10.1
10.2 | 162.7461 | 1.2719 | 2.8405 | 0.44779 | 2.9286 | 14.8997 | 10.2
10.3 | 163.7114 | 1.2884 | 2.8573 | 0.45092 | 2.9422 | 15.1631 | 10.3
10.4 | 164.6755 | 1.3049 | 2.8741 | 0.45402 | 2.9557 | 15.4278 | 10.4
10.5 | 165.6385 | 1.3214 | 2.8909 | 0.4571 | 2.969 | 15.6937 | 10.5
10.6 | 166.6005 | 1.338 | 2.9077 | 0.46015 | 2.9822 | 15.9608 | 10.6
10.7 | 167.5615 | 1.3546 | 2.9245 | 0.46317 | 2.9953 | 16.229 | 10.7
10.8 | 168.5217 | 1.3711 | 2.9413 | 0.46617 | 3.0082 | 16.4984 | 10.8
10.9 | 169.481 | 1.3877 | 2.958 | 0.46914 | 3.0209 | 16.7689 | 10.9
11 | 170.4396 | 1.4043 | 2.9747 | 0.47208 | 3.0336 | 17.0404 | 11
11.1 | 171.3976 | 1.4209 | 2.9915 | 0.475 | 3.046 | 17.3129 | 11.1
11.2 | 172.3549 | 1.4376 | 3.0082 | 0.47789 | 3.0584 | 17.5864 | 11.2
11.3 | 173.3117 | 1.4542 | 3.0249 | 0.48075 | 3.0706 | 17.8608 | 11.3
11.4 | 174.268 | 1.4708 | 3.0416 | 0.48358 | 3.0826 | 18.1362 | 11.4
11.5 | 175.224 | 1.4875 | 3.0582 | 0.48639 | 3.0945 | 18.4123 | 11.5
11.6 | 176.1796 | 1.5041 | 3.0749 | 0.48917 | 3.1063 | 18.6893 | 11.6
11.7 | 177.1349 | 1.5208 | 3.0916 | 0.49192 | 3.1179 | 18.9671 | 11.7
11.8 | 178.0901 | 1.5375 | 3.1083 | 0.49464 | 3.1294 | 19.2456 | 11.8
11.9 | 179.0451 | 1.5541 | 3.1249 | 0.49733 | 3.1408 | 19.5248 | 11.9
12 | 180 | 1.5708 | 3.1416 | 0.5 | 3.152 | 19.8047 | 12
12.1 | 179.0451 | 1.5875 | 3.1583 | 0.50264 | 3.1631 | 20.0852 | 12.1
12.2 | 178.0901 | 1.6041 | 3.1749 | 0.50525 | 3.174 | 20.3662 | 12.2
12.3 | 177.1349 | 1.6208 | 3.1916 | 0.50783 | 3.1848 | 20.6478 | 12.3
12.4 | 176.1796 | 1.6375 | 3.2083 | 0.51038 | 3.1955 | 20.9299 | 12.4
12.5 | 175.224 | 1.6541 | 3.225 | 0.51291 | 3.206 | 21.2125 | 12.5
12.6 | 174.268 | 1.6708 | 3.2416 | 0.51541 | 3.2164 | 21.4955 | 12.6
12.7 | 173.3117 | 1.6874 | 3.2583 | 0.51787 | 3.2267 | 21.7788 | 12.7
12.8 | 172.3549 | 1.704 | 3.275 | 0.52031 | 3.2368 | 22.0625 | 12.8
12.9 | 171.3976 | 1.7207 | 3.2917 | 0.52272 | 3.2468 | 22.3465 | 12.9
13 | 170.4396 | 1.7373 | 3.3085 | 0.5251 | 3.2566 | 22.6307 | 13
13.1 | 169.481 | 1.7539 | 3.3252 | 0.52745 | 3.2664 | 22.9151 | 13.1
13.2 | 168.5217 | 1.7705 | 3.3419 | 0.52977 | 3.2759 | 23.1997 | 13.2
13.3 | 167.5615 | 1.787 | 3.3587 | 0.53206 | 3.2854 | 23.4844 | 13.3
13.4 | 166.6005 | 1.8036 | 3.3755 | 0.53433 | 3.2947 | 23.7692 | 13.4
13.5 | 165.6385 | 1.8201 | 3.3922 | 0.53656 | 3.3039 | 24.054 | 13.5
13.6 | 164.6755 | 1.8367 | 3.4091 | 0.53876 | 3.3129 | 24.3388 | 13.6
13.7 | 163.7114 | 1.8532 | 3.4259 | 0.54093 | 3.3218 | 24.6235 | 13.7
13.8 | 162.7461 | 1.8697 | 3.4427 | 0.54308 | 3.3306 | 24.9081 | 13.8
13.9 | 161.7797 | 1.8861 | 3.4596 | 0.54519 | 3.3392 | 25.1926 | 13.9
14 | 160.8119 | 1.9026 | 3.4765 | 0.54727 | 3.3477 | 25.4769 | 14
14.1 | 159.8427 | 1.919 | 3.4934 | 0.54932 | 3.356 | 25.761 | 14.1
14.2 | 158.872 | 1.9354 | 3.5103 | 0.55134 | 3.3643 | 26.0447 | 14.2
14.3 | 157.8999 | 1.9518 | 3.5273 | 0.55333 | 3.3723 | 26.3282 | 14.3
14.4 | 156.9261 | 1.9681 | 3.5443 | 0.55529 | 3.3803 | 26.6112 | 14.4
14.5 | 155.9506 | 1.9844 | 3.5613 | 0.55722 | 3.3881 | 26.8939 | 14.5
14.6 | 154.9733 | 2.0007 | 3.5784 | 0.55911 | 3.3958 | 27.176 | 14.6
14.7 | 153.9942 | 2.017 | 3.5955 | 0.56097 | 3.4033 | 27.4577 | 14.7
14.8 | 153.0132 | 2.0332 | 3.6126 | 0.56281 | 3.4107 | 27.7388 | 14.8
14.9 | 152.0301 | 2.0494 | 3.6298 | 0.56461 | 3.418 | 28.0192 | 14.9
15 | 151.045 | 2.0655 | 3.647 | 0.56637 | 3.4251 | 28.299 | 15
15.1 | 150.0576 | 2.0817 | 3.6642 | 0.56811 | 3.4321 | 28.5781 | 15.1
15.2 | 149.068 | 2.0977 | 3.6815 | 0.56981 | 3.439 | 28.8564 | 15.2
15.3 | 148.076 | 2.1138 | 3.6988 | 0.57148 | 3.4457 | 29.1339 | 15.3
15.4 | 147.0815 | 2.1298 | 3.7161 | 0.57312 | 3.4523 | 29.4105 | 15.4
15.5 | 146.0845 | 2.1458 | 3.7335 | 0.57472 | 3.4587 | 29.6862 | 15.5
15.6 | 145.0848 | 2.1617 | 3.751 | 0.5763 | 3.465 | 29.961 | 15.6
15.7 | 144.0824 | 2.1775 | 3.7685 | 0.57783 | 3.4712 | 30.2347 | 15.7
15.8 | 143.0771 | 2.1934 | 3.786 | 0.57934 | 3.4772 | 30.5073 | 15.8
15.9 | 142.0689 | 2.2092 | 3.8036 | 0.58081 | 3.4831 | 30.7789 | 15.9
16 | 141.0576 | 2.2249 | 3.8213 | 0.58224 | 3.4888 | 31.0492 | 16
16.1 | 140.0431 | 2.2406 | 3.839 | 0.58364 | 3.4944 | 31.3183 | 16.1
16.2 | 139.0254 | 2.2562 | 3.8567 | 0.58501 | 3.4999 | 31.5861 | 16.2
16.3 | 138.0042 | 2.2718 | 3.8746 | 0.58634 | 3.5052 | 31.8525 | 16.3
16.4 | 136.9796 | 2.2873 | 3.8924 | 0.58764 | 3.5104 | 32.1176 | 16.4
16.5 | 135.9514 | 2.3028 | 3.9104 | 0.5889 | 3.5154 | 32.3812 | 16.5
16.6 | 134.9194 | 2.3182 | 3.9284 | 0.59013 | 3.5202 | 32.6433 | 16.6
16.7 | 133.8835 | 2.3336 | 3.9465 | 0.59132 | 3.525 | 32.9038 | 16.7
16.8 | 132.8436 | 2.3489 | 3.9646 | 0.59247 | 3.5296 | 33.1626 | 16.8
16.9 | 131.7996 | 2.3642 | 3.9828 | 0.59359 | 3.534 | 33.4198 | 16.9
17 | 130.7514 | 2.3793 | 4.0011 | 0.59467 | 3.5383 | 33.6752 | 17
17.1 | 129.6987 | 2.3945 | 4.0195 | 0.59571 | 3.5424 | 33.9288 | 17.1
17.2 | 128.6414 | 2.4095 | 4.038 | 0.59672 | 3.5464 | 34.1805 | 17.2
17.3 | 127.5795 | 2.4245 | 4.0565 | 0.59768 | 3.5502 | 34.4303 | 17.3
17.4 | 126.5126 | 2.4394 | 4.0751 | 0.59861 | 3.5539 | 34.6781 | 17.4
17.5 | 125.4408 | 2.4543 | 4.0938 | 0.59951 | 3.5575 | 34.9238 | 17.5
17.6 | 124.3637 | 2.469 | 4.1126 | 0.60036 | 3.5608 | 35.1674 | 17.6
17.7 | 123.2813 | 2.4838 | 4.1315 | 0.60117 | 3.564 | 35.4088 | 17.7
17.8 | 122.1933 | 2.4984 | 4.1505 | 0.60195 | 3.5671 | 35.6479 | 17.8
17.9 | 121.0996 | 2.5129 | 4.1696 | 0.60268 | 3.57 | 35.8847 | 17.9
18 | 120 | 2.5274 | 4.1888 | 0.60337 | 3.5727 | 36.1191 | 18
18.1 | 118.8942 | 2.5418 | 4.2081 | 0.60403 | 3.5753 | 36.351 | 18.1
18.2 | 117.7822 | 2.5561 | 4.2275 | 0.60464 | 3.5777 | 36.5804 | 18.2
18.3 | 116.6635 | 2.5703 | 4.247 | 0.60521 | 3.58 | 36.8071 | 18.3
18.4 | 115.5381 | 2.5845 | 4.2667 | 0.60574 | 3.5821 | 37.0312 | 18.4
18.5 | 114.4057 | 2.5985 | 4.2864 | 0.60622 | 3.584 | 37.2524 | 18.5
18.6 | 113.266 | 2.6125 | 4.3063 | 0.60667 | 3.5857 | 37.4709 | 18.6
18.7 | 112.1188 | 2.6264 | 4.3263 | 0.60707 | 3.5873 | 37.6864 | 18.7
18.8 | 110.9638 | 2.6402 | 4.3465 | 0.60742 | 3.5887 | 37.8989 | 18.8
18.9 | 109.8007 | 2.6538 | 4.3668 | 0.60773 | 3.5899 | 38.1083 | 18.9
19 | 108.6293 | 2.6674 | 4.3872 | 0.608 | 3.591 | 38.3145 | 19
19.1 | 107.4493 | 2.6809 | 4.4078 | 0.60821 | 3.5918 | 38.5175 | 19.1
19.2 | 106.2602 | 2.6943 | 4.4286 | 0.60839 | 3.5925 | 38.7171 | 19.2
19.3 | 105.0618 | 2.7076 | 4.4495 | 0.60851 | 3.593 | 38.9133 | 19.3
19.4 | 103.8538 | 2.7208 | 4.4706 | 0.60859 | 3.5933 | 39.1059 | 19.4
19.5 | 102.6356 | 2.7338 | 4.4919 | 0.60862 | 3.5934 | 39.2949 | 19.5
19.6 | 101.407 | 2.7468 | 4.5133 | 0.6086 | 3.5933 | 39.4802 | 19.6
19.7 | 100.1676 | 2.7596 | 4.5349 | 0.60852 | 3.593 | 39.6616 | 19.7
19.8 | 98.9168 | 2.7723 | 4.5568 | 0.6084 | 3.5926 | 39.8392 | 19.8
19.9 | 97.6542 | 2.7849 | 4.5788 | 0.60823 | 3.5919 | 40.0126 | 19.9
20 | 96.3794 | 2.7974 | 4.601 | 0.608 | 3.591 | 40.1819 | 20
20.1 | 95.0917 | 2.8098 | 4.6235 | 0.60772 | 3.5899 | 40.347 | 20.1
20.2 | 93.7907 | 2.822 | 4.6462 | 0.60738 | 3.5885 | 40.5077 | 20.2
20.3 | 92.4756 | 2.8341 | 4.6692 | 0.60699 | 3.587 | 40.6638 | 20.3
20.4 | 91.146 | 2.8461 | 4.6924 | 0.60653 | 3.5852 | 40.8153 | 20.4
20.5 | 89.8011 | 2.8579 | 4.7159 | 0.60602 | 3.5832 | 40.9621 | 20.5
20.6 | 88.4401 | 2.8696 | 4.7396 | 0.60545 | 3.581 | 41.1039 | 20.6
20.7 | 87.0623 | 2.8812 | 4.7637 | 0.60482 | 3.5785 | 41.2407 | 20.7
20.8 | 85.6669 | 2.8926 | 4.788 | 0.60413 | 3.5757 | 41.3722 | 20.8
20.9 | 84.2528 | 2.9038 | 4.8127 | 0.60337 | 3.5727 | 41.4984 | 20.9
21 | 82.8192 | 2.9149 | 4.8377 | 0.60254 | 3.5695 | 41.619 | 21
21.1 | 81.365 | 2.9259 | 4.8631 | 0.60165 | 3.5659 | 41.734 | 21.1
21.2 | 79.889 | 2.9367 | 4.8889 | 0.60068 | 3.5621 | 41.843 | 21.2
21.3 | 78.3899 | 2.9473 | 4.915 | 0.59965 | 3.558 | 41.9459 | 21.3
21.4 | 76.8664 | 2.9577 | 4.9416 | 0.59853 | 3.5536 | 42.0425 | 21.4
21.5 | 75.3169 | 2.968 | 4.9687 | 0.59734 | 3.5489 | 42.1325 | 21.5
21.6 | 73.7398 | 2.9781 | 4.9962 | 0.59607 | 3.5439 | 42.2158 | 21.6
21.7 | 72.1332 | 2.988 | 5.0242 | 0.59472 | 3.5385 | 42.292 | 21.7
21.8 | 70.495 | 2.9977 | 5.0528 | 0.59328 | 3.5328 | 42.3609 | 21.8
21.9 | 68.823 | 3.0072 | 5.082 | 0.59174 | 3.5267 | 42.4221 | 21.9
22 | 67.1146 | 3.0165 | 5.1118 | 0.59011 | 3.5202 | 42.4754 | 22
22.1 | 65.3669 | 3.0257 | 5.1423 | 0.58838 | 3.5133 | 42.5204 | 22.1
22.2 | 63.5767 | 3.0345 | 5.1736 | 0.58655 | 3.506 | 42.5566 | 22.2
22.3 | 61.7401 | 3.0432 | 5.2056 | 0.5846 | 3.4982 | 42.5836 | 22.3
22.4 | 59.8529 | 3.0516 | 5.2386 | 0.58254 | 3.49 | 42.601 | 22.4
22.5 | 57.91 | 3.0598 | 5.2725 | 0.58034 | 3.4812 | 42.6081 | 22.5
22.6 | 55.9058 | 3.0678 | 5.3074 | 0.57801 | 3.4719 | 42.6043 | 22.6
22.7 | 53.8331 | 3.0755 | 5.3436 | 0.57554 | 3.462 | 42.5889 | 22.7
22.8 | 51.6839 | 3.0829 | 5.3811 | 0.5729 | 3.4514 | 42.561 | 22.8
22.9 | 49.4479 | 3.09 | 5.4202 | 0.57009 | 3.4401 | 42.5196 | 22.9
23 | 47.1129 | 3.0968 | 5.4609 | 0.56709 | 3.428 | 42.4635 | 23
23.1 | 44.6633 | 3.1033 | 5.5037 | 0.56386 | 3.415 | 42.3911 | 23.1
23.2 | 42.0789 | 3.1095 | 5.5488 | 0.56039 | 3.401 | 42.3006 | 23.2
23.3 | 39.3333 | 3.1153 | 5.5967 | 0.55663 | 3.3857 | 42.1896 | 23.3
23.4 | 36.3897 | 3.1207 | 5.6481 | 0.55252 | 3.3691 | 42.0548 | 23.4
23.5 | 33.1957 | 3.1257 | 5.7038 | 0.54799 | 3.3506 | 41.8917 | 23.5
23.6 | 29.6702 | 3.1302 | 5.7653 | 0.54293 | 3.33 | 41.6933 | 23.6
23.7 | 25.6771 | 3.1342 | 5.835 | 0.53713 | 3.3062 | 41.4486 | 23.7
23.8 | 20.9506 | 3.1375 | 5.9175 | 0.53021 | 3.2777 | 41.1363 | 23.8
23.9 | 14.804 | 3.1402 | 6.0248 | 0.52121 | 3.2405 | 40.703 | 23.9
23.95 | 10.4644 | 3.1411 | 6.1005 | 0.51489 | 3.2143 | 40.3852 | 23.95

## ESRI_MAPINFO_SHEET (hidden)
DO NOT EDIT
For Esri use only